Valve controller documentation draft 26.12.2014

# Controller board pins assignment
The valve controller is operated by a TI MSP430 LaunchPad development board with msp430g2553 controller.
Controller pins are assigned the following way:
Pin	Function	Direction	States
VCC	3.3 V power	IN	N/A
P1.0	N/C (Integrated red LED)	N/C	N/A
P1.1	N/C (Integrated UART)	N/C	N/A
P1.2	N/C (Integrated UART)	N/C	N/A
P1.3	N/C (Integrated button)	N/C	N/A
P1.4	ADC Vref+	IN	Vref/5 analog signal (around 2V DC)
P1.5	ADC Ch.5	IN	Vtrig/5 analog signal
P1.6	N/C (Integrated green LED)	N/C	N/A
P1.7	ADC Ch.7	IN	Vin/5 analog signal
P2.0	MCU mode out, / activated by internal switch	OUT	High: output enabled, valve open / Low: output disabled, valve closed
P2.1	Mixed mode close	OUT	High: normal operation / Low: short pulse to activate the trigger and disable the output(close the valve)
P2.2	N/C	N/C	N/A
P2.3	Mixed mode open	OUT	High: short pulse to activate the trigger and enable the output (open the valve) / Low: normal operation
P2.4	Valve state indicator	IN	MCU internal pull-up to high, / LOW: valve closed
P2.5	Valve state indicator	IN	MCU internal pull-up to high, / LOW: valve open
GND	Board ground level	IN	N/A

# Serial interface
Valve controller emulates a serial port and is operated using a simple serial-based protocol.
## Serial port settings
9600 bit/s, 8 bytes, 1 stop bit, no parity control.
Setting of any bitrate higher than 9600 is internally overriden to 9600 bit/s by the USB-Serial converter.
## Communication protocol
Valve controller supports a limited set of single-character commands. There is no need to follow the command by CR-LF sequence. Controller immediately responds with a sequence of lines, each terminated by CR-LF sequence.
A summary of commands is provided in the following table:
Command	Symbol	Detailed description
Identify	i	Print the board identification string.
Get Voltages	v	Print actual input voltage and corresponding pressure. / Response format is the following: / Vin: 57mV / Pin: 1.4E-11Torr
Get valve position	p	Get the valve position indicator status. Possible answers are: / UNDEF Both position switches read as open or closed. / OPEN Position switch “open” reads as 0, “close” as 1 / CLOSED Position switch “open” reads as 1, “close” as 0
Close the valve	c	Immediately send a trigger pulse to close the valve, / deactivate the digital output. Returns “OK” as a confirmation
Open the valve	o	Check that Vin<Vthresh (both hardware and software), send a trigger pulse to open the valve and activate the output if true. / Returns “OK” if the operation succeded. / Returns “PERR” if Vin>=Vthresh.
Get software threshold	s	Get the software-defined threshold. / Response format is the following: / Vst: 1234mV / Pst: 2.4E-10Torr
Get hardware threshold	t	Get the hardware-defined threshold voltage and corresponding pressure. Response format is the following: / Vtr: 1234mV / Ptr: 2.4E-10Torr
Unknown command	any	Board responds with a string “Unknown command” to any invalid symbol.

# Operation
Valve controller can operate independently or in junction with a personal computer.
## Independent operation
In independent mode controller constantly tracks the input voltage and cuts the output at the moment the input voltage becomes higher than either software or hardware defined threshold voltages.
Hardware-defined threshold voltage is adjusted by a variable resistor and can be monitored in PC-conneced mode using the terminal command “v”.
Software-defined threshold is set using a short press of the “set” button: valve controller remembers the current input voltage as a threshold.
To clear the threshold voltage, a “set” button should be pressed for more than a second.
## Connected operation
In connected mode controller works exactly the same way as in independent mode, plus all of the terminal commands described before become available.